## Slide 1: Operation and Upgrade of the SLS RF Systems
Lukas Stingelin :: RF-Systems 1 :: Paul Scherrer Institute
25th European Synchrotron Light Source RF (ESLS-RF) workshop organized by DESY November 2021

## Slide 2: SLS-Operation Statistics until 2020
2020:
During Corona lock-down, slightly reduced User Operation
Good MTBF, but lowest value of time between distortions (16h)!
34% of down-time due to RF
Outages more than 5h:
Broken LINAC RF window (26h)
Water leaks SR cavity (5h)
Several Interlocks due to bad Klystron filament connector

## Slide 3: 2021 SLS Failures
Obstacle (diagnostics mirror) in the LINAC beam path (not in control system)
Power glitches, thunderstorms, power outage due to short in 16kV distribution
Refill of demineralized water caused flow interlock at RF stations
Interlocks of RF-stations due to coupler valve error (pressurized air, or PLC error?)
LINAC Klystron failure
LINAC local structure cooling unit broken
Failure of reflected power measurement in Booster SSA
Failure of 6-way splitter in the Booster SSA
Cooling Rack Temperatur controller got unstable
So far: 10% due to RF

## Slide 4: Arcing at LINAC Tank HV transformer Part 2
In 2020: Installed isolation plate:
In Mai 2021: EOLC (+VSWR)-Interlocks: Black marks at plate. Replaced HV-Transformer
In July 2021: Again EOLC-Interlocks: Replaced Klystron

## Slide 5: Water Leaks in the SR ELETTRA cavities
3 water leaks in the cavity wall-cooling circuit in 2020 (old cavities had problems with the flange cooling circuit, improved in the new cavities)
Improved water chemistry
Demiwater instead of tap water
Increased pH to 10
Replaced Polyflex by:
  NALCO TRAC100 (Molybdate)
  NALCO 73199 (Azole)
  NALCO Biocide 733O
Upgrade of Cooling Racks (see following slides)
Courtesy: SvoBaTech
 Pressure (here after the pump) is not an issue.

## Slide 6: Upgrade 1: New sample point
New sample point on the highest point of the circuit
Defined method to take samples:
The sampling valve should be quickly opened and closed three times.
Leave the water run with a constant flow of 100ml to 200ml per minute for about 3–5 minutes.
Take the samples without touching or closing the valve before the samples are taken.
Courtesy: D. Kunz

## Slide 7: Upgrade 2: Viewing Window
Recommended by SvoBaTech:
To see gas-bubbles in circulation
To see copper-, iron-dust, or sludge
To observe color-change through the window during the cleaning process.
Courtesy: D. Kunz

## Slide 8: Upgrade 3: New filtering
The small bypassfilter was removed and replaced by a full stream filter.
The new filter contents a filter cartridge.
If there is a large difference between the inlet and outlet pressure the cartridge has to be replaced.
Courtesy: D. Kunz

## Slide 9: Current Problems 1: High Bacterial Count on one station
SR4: Very high bacterial count since march 2021. The CLRK was cleaned with NALCO 2567 and four times flushed with demineralized water  Situation improved
Courtesy: D. Kunz
Regular check of the water quality, for example:
Conductivity
Acidity
Copper and iron
Azoles
Bacterial count
Oxygen saturation

## Slide 10: Current Problems 2: Water loss
SR1: On CLRK of SR1 we lose continuously water. Probably in the heat exchanger between primary and secondary circuit. Repair will be done in January Shutdown 2022, by exchange of the gaskets.
SR3: On CLRK of SR3 we lose water when the heat exchanger is cold. Repair will be done in January Shutdown as well.

## Slide 11: SLS RF works
Done:
Improvements on «Pantak» connectors
Installation of filters etc. to the cavity cooling rack
Removal of «3GHz Teststand» in SLS LINAC
Upgrade of Super-3HC vacuum controller
Installation of additional crane for SLS 2 installation
In Progress:
LINAC pulser board upgrade to shorter bunches
LINAC timing fine-shift according to bucket-number
LINAC: HV-Pulse recording
Renewal of water cooling piping
Storage-Ring+Booster cavities: HOMFS upgrade
Storage-Ring: LLRF upgrade
Modification of 500MHz Teststand
SLS2 RF Upgrade  See next slides

## Slide 12: SLS 2.0 Master Milestones (version Oct. 2021)
5
start dark time
02.10.2023
6
tunnel closure
20.12.2024
7
nominal beam available
01.05.2025
8
start user operation, BL-Grp.1
01.08.2025
Now: 8.11.2021
1
storage ring lattice defined
30.06.2020
2
beamlines & beamline positions defined
30.09.2020
2020 | 2021 | 2022 | 2023 | 2024 | 2025 | 2026
9
start shutdown phase 2
21.12.2025
10
re-start user operation
01.06.2026
3
SLS2.0 funding secured
01.01.2021
Infrastructure not available
Kick-Off Meetings with RI and Cryoelectra
4
ready for DT
30.09.2022
Teststand bunker modified for new cavities
Page 12

## Slide 13: SLS 2.0 Main RF parameters
Main RF-System | SLS 2.0 |  | SLS
Total voltage [kV] | 1440 | 1780 | 2080
Energy acceptance (without harmonic cavity) | 5% → 3.7% | 6.3% → 5.2% | 3%
Number of cavities | 4 |  | 4
Voltage per cavity [kV] | 360 | 445 | 520
Wall loss per cavity [kW] | 20 | 30 | 40
Required RF-power with beam and minimum ID Power [kW] | 88 | 98.1 | 95
Required RF-power with beam and maximum ID Power [kW] | 114 | 124 | 100
Optimal coupling | 4.62…6.0 | 3.4…4.3 | 2.5
Detuning for matching [kHz] | -49.1…-57.1 | -32.4…-44.6 | -33
HOM control | By strong HOM damping |  | Temp. detuning
Amplifier Spec: 150kW (new solid state amplifiers)
 For limited emergency operation with 3 stations only, coupler must work also at 150kW!
Main RF-System | SLS 2.0
Arcing rate at Max. operating voltage | <0.5/month
Cooling water properties | Demineralized water, pH neutral, oxygen depleted
Max. relative water pressure | 10 bar for test, 6 bar operational
Max. water speed in the cooling channels | ≤ 1.5 m/s
Max. water flow of the entire Cavity assembly | ≤140 l/min
Max. inlet – outlet pressure drop | ≤ 4 bar

## Slide 14: RF Power Considerations for 4 Cavity Operation
(Courtesy: R. Kalt)
Static, without consideration of power limiter:
Goal: Protect the cavity and coupler, but at a sudden beam loss do not block the high-power RF operation  fast re-start of beam accumulation.
Expected LLRF pwr. limiter controller reaction is faster than 20us pwr. error below 1%.
Coordination of RF-ILK and LLRF expert settings required.

## Slide 15: Project SLS 2.0: Cavity Procurement
Page 15
HOM damper
Tuner
Fundamental power coupler and transition for 150kW
Ordered 5 HOM damped cavities from Research Instruments
Coupler should handle 150kW input power:
Evaluate PETRA-Type coupler, like DESY
Cavities will be connected to the demineralized water circuit of the storage ring. (No cooling rack anymore!), Air cooling from pressurized air distr.
Flow velocity < 1.5m/m : Hopefully no water leaks
Cavities to be mounted in the teststand and in the storage-ring: Additional supports required.
Water cooling for HOM ferrites

## Slide 16: Preliminary Layout: 500MHz RF-System
All main cavities in the same long straight X05
Modification of tunnel wall to host cavities and neighboring beam line.
Connection via Waveguide – Coax transition to the cavity coupler. (WaTraX, Doorknob?)
Next to the cavities, there will be collimators
Collimators (under the Quads)

## Slide 17: 500MHz Cavities in Long Straight X05: Vacuum
Preliminary draft layout for the case of HOM-damped cavities:
  Taper at the beginning and end of the 4 cavities
  Vacuum pumps (IGP+ZAO) and -sensors at both sides of the cavities
  Valves and bellows to separate the cavities for installation and conditioning.
Characterization of the cavities in a clean-room, high power tests + conditioning in the teststand
  485 mm upstream and downstream at each cavity: pump cross with:
  2x 300 l/s ion getter pump
  +1x 1000 l/s NEG pump
  1 vacuum probe on the pumping cross
  1 vacuum probe on each cavity
  View-port on coupler (like at DESY)
  Coupler and Beam-pipe flanges to be defined

## Slide 18: Investigation of Synchrotron Radiation on Cavity Surfaces (Synrad Calculation by Vacuum Section)
There is no direct synchrotron light hitting the cavity surfaces
But indirect, from reflection on the walls of the beam-pipe, there will be some light on the cavity.
What is the limit, how does it affect the arcing rate?
Radiation spectrum; X axis in eV Y axis in photons/ sec/ 0.1%BW.
Radiation distribution at the cavity.

## Slide 19: 500MHz Amplifier Procurement
Signed contract with Cryoelectra for 4 high power solid state UHF 150kW amplifiers + spares
Approved ProkW funding proposal for energy saving (hope for efficiency > 56%).
It is planned to add an external high power circulator
 Layout with RF platform
 Earthquake safety investigations ongoing
 Water cooling concept is in progress.

## Slide 20: Thanks for your attention!
Questions?